- Objective

To lay down a procedure for Operation, Cleaning and Calibration of weighing Balances.

- Scope

The SOP is applicable for Operation, Cleaning and Calibration of weighing Balances of Alkem Laboratories Ltd., Unit-II, Baddi.

- Responsibility

Production Technician: To carry out the activity in accordance with the procedure as mentioned in the SOP.

Production Officer/ Executive: To ensure that the activity is carried out in accordance with the procedure as mentioned in this SOP.

- Accountability

Head Production.

- Procedure

5.1       Pan Balance

5.1.1   Operation / Cleaning

    - Clean the weighing chamber, weighing pan with a clean and dry lint free cloth.

5.1.3  Check that the power and printer cables are properly connected to the balance.

    - Check the levelling of the balance with bubble position in centre and adjust if required.

    - Turn ‘ON’ the instrument and leave the balance to automatic self test.

    - At the end of the self-test, a zero read out is displayed. This means that the                                                              balance is ready to Operate.

  - Daily  Verification

    - Tare the balance.

    - Verification of balance in respective area shall be done by standard weights mentioned in               Annexure – III,   and put the weights with the help of the forceps on the pan.

    - Close the cover shield and wait till the display shows the stable reading and record the   readings in Annexure - I.

  - Verification frequency

    - Check and record daily before start of day the activity.

    - Check and record after resuming of power supply (in case of power failure).

    - Check and record after any maintenance work.

Note:  Recording of the zero reading shall be done as per the actual display of the weighing scale.

  - Monthly Calibration

    - Uncertainty

    - Use 100 mg standard calibrated weight for measurement of uncertainty.

    - Place the weight on the pan.

    - Close the shield cover.

    - Wait for display to stabilize.

    - Record the reading in Annexure –II.

    - Remove the weight and again place the weight on the pan.

    - Repeat the above steps to take ten consecutive readings for the same weight.

    - Calculate the standard deviation of ten readings & Uncertainty.

Random error + Systematic error

Measurement of Uncertainty = ------------------------------------------ 	x 100

Actual weight

Random error = 3 x Standard Deviation (n=10)

Systematic error = Observed weight (Average) - Actual weight

Acceptance Criteria: Not more than 0.1%

- Drift

2

5        1       3

4

4

1      2     3

5

Take a standard weight of 100 mg and place it on five different positions of the Pan. Select   positions according to shape of the pan as shown below. Record the weight observed on the five different places in Annexure-II.

1		2

3

4                               5		5

- Calculate the mean of five readings. None of the individual weight should be out side ± 0.2 mg of that average value.

- Put the calibration status label on the instrument.

- After calibration, put the “CALIBRATION TAG” on the balances.

6.1 Platform balance

6.1.1 Procedure

6.1.1.1 Zero Reading

6.1.1.2 Zero reading of the weighing scales is one of the indications of the proper functioning of the scale, after every switch ‘ON’ activity.

6.1.1.3 Zero reading of the weighing scale must be checked and recorded before use to avoid erroneous results. Record as it displays on the balance screen.

6.1.1.4 The errors in the zero reading may occur due to:

- The weighing scales not being in level as per the water level.

- Due to mechanical or electronic defect.

- The platform not being clean.

- Some extraneous matter hampering the smooth operation of the platform load cell.

6.1.1.5 Procedure to avoid zero reading error:

- Clean the platform of the weighing scale with a clean lint free dry cloth.

- Check that the weighing scale is in level, if not adjust with the help of adjustable knobs.

- Switch on the weighing scale.

- Check for the Zero reading. If there is any error, rectify the error in the zero reading using ‘zero’ key or refer for maintenance.

6.1.1.6 Verification Frequency

- Check and record daily before start of day the activity.

- Check and record after resuming of power supply (in case of power failure).

- Check and record after any maintenance work.

Note: Recording of the zero reading shall be done as per the actual display of the weighing scale.

6.1.1.7 Procedure for Daily / Monthly verification of platform balances

- Switch “ON” the weighing scale.

- Ensure the display for “zero” reading. Rectify if necessary.

- Get the standard weights.

- Ensure that the standard weights are handled with gloved hands.

- For Daily verification select the minimum and maximum standard weights as per Annexure –III

Of respective area and keep at the middle of the platform.

- For Monthly verification keep the standard weights on platform at five different locations as per below given location plan. Use minimum, middle and maximum standard weights for monthly verification as per Annexure – III, of respective area.

1		2

3

4                           5     55                                 		5

- Record the displayed readings of all above mentioned locations in Annexure – I. It should be within specified tolerance limits mentioned in Annexure –III of respective area.

- If the variation is within the specified tolerance limit then continue to use the weighing scale.

6.1.1.8 If the variation is not within the specified tolerance limits

- Label the weighing scale with a tag, “UNDER MAINTAINANCE”.

- Inform the Department Head and Engineering Department.

- Do not use weighing scale until all errors are rectified. After rectification, verify the weighing scale before using it. Record the verification details in respective format.

- If any balance is found erroneous, review usage of balance since last verification and intimate it to QA department.

  - Annual calibration

6.2.1  Platform balances should be calibrated once in a year by Department of Weights and measurements Government of India.

    - Update below given Calibration status label after performing calibration.

| BALANCE ID |  |
| --- | --- |
| Calibrated By |  |
| Date Of Calibration |  |
| Due Date Of Calibration |  |
| Calibration Status Checked By / Date |  |

6.3 Calibration of Standard weights

6.3.1 Standard weights should be calibrated once in a year by Department of Weights and measurements                    Government of India.

6.3.2 If any change in the location of the balances or removal/addition, the same shall be handled     through SOP CQA/032 “Change Management System”.

6.3.3 Removal of balance from the facility shall be handled through SOP QA/024 “Procedure for Disposal                  of Rejected /Waste Materials and Samples” and filled destruction note shall be attached with master SOP.

7.0 Annexure (S)

Annexure - I 	:	Daily Verification Record of Pan / Platform Weighing Balances.

Annexure - II 	:	Monthly calibration Record of Pan / Platform Weighing Balances.

Annexure – III: 	Verification strategy For Weighing Balances.

8.0 Reference (S)

NA

9.0 Glossary (S):

SOP		:	Standard Operating Procedure

No.		:	Number

IPQA		:	In process quality Assurance

Std wt.	:	Standard weight

Obs wt.	:	Observed weight

10.0 Revision History

| Revision No. | Details Of Changes | Reason for Change |
| --- | --- | --- |
| 00 | New SOP | NA |