- Revision History

| Rev.No. | Details of changes | Reason for change |
| --- | --- | --- |
| 01 | NA | Periodic Review |
| 02 | NA | Periodic Review |

2.0    Objective

To lay down a procedure for Operation & Cleaning of Filter Integrity test Apparatus.

3.0    Scopes

This SOP is applicable for the Operation & Cleaning of Filter Integrity test Apparatus of C Block.

# 4.0   Responsibility

Technician			:	Operation and Cleaning.

Officer and Above Production	:	Implemented the Procedure.

5.0	   Accountability

Production Head.

- Procedure

6.1    Start Up

6.1.1 Before starting the Filter Integrity test Apparatus ensure that the compressed air supply is available.

6.1.2 Ensure that the sterilized filtration assembly is available for bubble point test.

6.1.3 Ensure that the filter is available on the work place for integrity testing.

6.1.4 Ensure that the filter is properly wetted in required wetting liquid for integrity testing.

6.1.5 Ensure that the power supply is available.

  - Operation

6.2.1 Connect the Palltronic filter integrity testing apparatus to compressed air supply with the help of festo tube joint with inlet valve (pressure of compressed air up to 3 kg/cm2).

6.2.2 Now connect the output to inlet of the exhaust vent valve and outlet of the valve to inlet of filter holder.

6.2.3  Connect the Palltronic filter integrity testing apparatus to power supply and switch ON the main.

6.2.4  Apparatus will indicate its identity and enter in to self-test.

6.2.5  Allow it to completion of self-test before starting the testing operation.

6.2.6 Press the air clean and press start allow the apparatus to air clean itself.

6.2.7 MMI of apparatus will display the following menus: -

TEST PROGRAMMES

FORWARD FLOW

TEST RESULTS

WATER INSTRUSION

AIR CLEAN

BUBBLE POINT

FLOW CHECK

FF & BP

SYSTEM

6.2.8  To run the bubble point test, press the bubble point on the main menu screen, screen will show        bubble point test, press input on test bar.

6.2.9 Feed the following details: -

- Operator.

- Production area.

- Filter line.

- Product name.

- Product batch no.

- Filter port no.

- Filter serial no.

- Housing.

- Wetting liquid.

- Test gas.

- BP minimum.

6.2.10 After making correct entries, press ok.

6.2.11 Screen will show bubble point test.

6.2.12 Press start to run the test.

6.2.13 After completion of the test, the apparatus will indicate the test result and print out of same will be printed.

6.2.14 Similarly run the forward flow and water instrusion as per filter type (for hydrophilic filter-forward flow and for hydrophobic filter- water instrusion).

7.0      Annexure (S)

NA

8.0     Reference (S)

NA

9.0     Glossary

SOP   	: Standard Operating Procedure.

NO      	: Number

PC		: Production C-Block