- Objective

Procedure For Indent, Receipt, Storage and Handling of Sieves and Screens.

- Scope

This SOP is applicable to sieves and screens used for vibratory sifter, multimill in the production areas of Alkem Laboratories Unit-II.

- Responsibility

Production Officer/ Executive		:	Implementation of procedure

Documentation and data entry.

Operator				:	Execution of operation.

- Accountability

Production Head			:	System compliance

- Procedure

  - Indenting

    - The codification for sieve and screens shall be decided at the time of indent, so That party should provide the new sieve and screens with proposed code number. Party should also provide the validation certificate and lead free certificate for Sieves/Screen and non toxicology test report of silicone gasket for sieves.

  - Coding System for Sieves and Screens

    - The coding system for sieves shall be as given below

Where “x” indicates Alkem laboratories Unit-II, 30 as sieve size, 001 as sieve Number & 08 as year of purchase respectively.

    - The coding system for Screen shall be as given below:

IS/2.5/001/08

Where “IS” for screens of Immunosuppressant block, 2.5 as screen size, 001 as

Screen number & 08 as year of purchase respectively. Check the sieve/screen stock Record (Annexure-VII) and take the new Sieve/screen Code No. which will be followed by last Code No. of the Sieve/Screen. Fill up the Sieve/Screen Indent Record (Annexure-VIII) with required Sieve/Screen Code No., and required certificates, then send to the supplier along with Regular SAP/Manual Indent.

  - On receipt of new sieve check the following:

    - Check the validation certificate, lead free certificate and non toxicology test report of silicone gasket.

    - Diameter-Check with scale.

    - Physical appearance-Check with naked eye.

    - Integrity-Check with naked eye/Integrity test microscope.

    - Mesh Size-Check with Normal Aperture Calibrator.

    - Coding Details-Check as per mentioned in the Sieve/Screen Indent Record. (Annexure-VIII).

  - Mesh size Checking Procedure.

    - For Mesh Size 4#to 40#- the distance between adjacent parallel wires is called as Nominal Aperture size (NAS). Nominal Aperture Calibrator is used to measure the NAS. The Nominal Aperture Calibrator has an illuminated scale with a full-scale gradation of 10mm divided in 100 parts. The least count of the scale is 0.1mm or 100 microns.

    - Take measurement as follows:

      - The calibrator is placed on the wire mesh & the individual opening is visualized by using the fine-tuning.

      - The distance between adjacent parallel wires is measured in both wrap & weft directions.

      - Nine such readings are taken all over the mesh.

      - The average of these readings is the Nominal Aperture Size of the wire mesh.

      - The NAS is compared with the standard sizes specified in the Annexure-IX.

      - If it falls within the specified tolerances the sieves is accepted.

      - After compliance of new sieve enters details in New sieve check Record Annexure-IV.

    - For Mesh Sizes Above 40# up to 200# Take measurement as follows:

      - The Calibrator is place on the wire mesh & the no of wires in 2.5mm is counted.

      - Nine such readings are taken all over the mesh.

      - Multiplying this by 10 will give the no wires in a linear inch, i.e 25.4mm approx.

      - The wire diameter of a particular mesh is taken from Annexure-IX.

      - The NAS of the wire mesh is calculated by using the below formula.

      - NAS= [(25.4X1000)/N]-D, Where N=Number of wires per Linear inch & D=Diameter for wire in microns.

      - The NAS is compared with the standard sizes specified in Tolerances on Aperture Widths for Sieves. Annexure-IX.

      - If it falls within the specified tolerances the sieves is accepted.

      - After compliance of new sieve enter details in new sieve check Record Annexure-I.

  - On receipt of new screens check the following:

    - Diameter-Check with scale.

    - Physical Appearance-check with naked eye.

    - Integrity-check with naked eye/Integrity test microscope.

    - Pore Size-Check with Nominal Aperture Calibrator.

    - Welding Quality-Check with naked eye.

    - Coding Details-Check as mentioned in the Sieve/Screen Indent Record. (Annexure-VIII).

  - Screen Size Checking Procedure:

    - The Calibrator is placed on the screen & the individual pores are visualized by using the fine-tuning.

    - Measure the diameter of the pores at 10 different locations on screen.

    - The average of these readings in Nominal Aperture Size (NAS) of the screen.

    - The NAS is compared with the standard sizes specified in the Annexure-X.

    - If it falls within the specified tolerances the screen is accepted.

    - After compliance of new screen enter details in Annexure-V. for multimill screen and Annexure-VI. For Cone Mill Screen.

  - Cleaning procedure:

    - Take used sieve/screen to adjacent washroom, in case where washroom facility is not available then pack in a poly bag and transfer if to IPC washroom area for cleaning.

    - Wash the sieve/screen with potable water to remove adhered powder.

    - Scrub the sieve/screen using nylon brush and wash thoroughly with potable water.

    - Then finally rinse with purified water.

    - Dry the sieve/screen with compressed air by applying low pressure.

    - Check the cleanliness and integrity of dried sieve/screen, if cleaning is not satisfactory reclean the sieve/screen.

    - Pack the cleaned, dry, inspected sieve/screen in poly bag and take in “change part room” make Entry in Issuance, usage, cleaning an retrieval record of sieves/screens Annexure-II.

  - Storage, Handling and Ussage.

    - Check for cleanliness and integrity, while receiving and issuance of sieve/screen.

    - Store the cleaned, polythene wrapped, labeled sieve/screen in the

    - Designated cabinet in CHANGE PARTS ROOM, Provided for different type Of sieves, screens.

    - Record the integrity of sieve/screen before and after use in Batch Manufacturing record.

    - Take the required size sieve/screen from the designated cabinet and enter In “Issuance” Usage, cleaning and Retrieval Record of Sieves/Screens Annexure-II. Also make the entry for cleaning and returning. Update Sieve/screen stock record (Annexure-VIII) whenever changes in Quantity of sieves/screens.

  - Destruction

    - In case any damage on sieve/screen the same shall be separated and Labelled as rejected.

    - Production Officer/Executive shall check the rejected Sieve/screen & Verified with QA.

    - Quarantine the rejected Sieve/Screen and raise destruction note collected From QA Dept. After authorization of destruction form, destruct rejected And quarantined Sieve/screen by defacing.

    - Fill the Sieve/Screen Destruction Record (Annexure-III) mentioning the Sieve/screen Code No. Date of Destruction, Destruction note No., and Reason for Destruction.

  - Annexure (S)

Annexure-I	New sieve check Record for sieve Size above #40 Up to#20.

Annexure-II	issuance, usage, cleaning and retrieval record of sieves and screens

Annexure-III	Sieve/screen Destruction record.

Annexure-IV	new sieve check record for sieve size #4 to #40.

Annexure-V	New Screen checks record for multi mill.

Annexure-VI	New screen check record for cone mill.

Annexure-VII	Sieve/screen stock record.

Annexure-VIII	Sieve/screen indents record.

Annexure-IX	Tolerance on Aperture widths for sieves.

Annexure-X	Tolerance on Aperture widths for sieves.

- Reference (S)

Sieves/screen calibration kit manual by pharma spares.

- Glossary

SOP		:	Standard Operating Procedure

NO		:	Number

- Annexure (S)

NA

- Reference (S)

SOP No. GN/012.

- Glossary

SOP		:		Standard Operating Procedure

NA		:		General

IPA		:		Isopropyle alcohol

QTY		:		Quantity