- Objective

To lay down a procedure for Calibration, Operation and Cleaning of weighing Balances.

- Scope

The SOP is applicable for Calibration, Operation and Cleaning of weighing Balances of Alkem Laboratories Unit-II.

- Responsibility

Officer/ Executive: Concerned Department

- Accountability

Head Production.

- CALIBRATION:

    - Before calibrating the balances, clean the weighing platforms and pans.

    - Adjust spirit level of balance to Zero, with the help of the spirit level/ leveling screws of the balance.

    - Switch on the mains

    - Press” TARE” continuously till Display shows 1/3 capacity of balance.

    - Then put the 1/3 weight of total capacity of balance, display starts blinking.

    - Again press “TARE” till blinking stops

    - As per “FIVE POINT CALIBRATION” Place the same weight on different five positions on   platform and check the reading.

x x x x x

    - Now balance is Calibrated and ready for “LINEARITY” check

    - Select the Standard Weights for Linearity of various Balances.

Selection of the Standard weights for various Balances.

| Balance Capacity | Standard Weights to be used | Least Count in kg/Gm |
| --- | --- | --- |
| 6.0 Kg | 1KG,2KG,5KG | 0.0005 kg |
| 15.0 Kg | 2KG,5KG,10KG | 0.001kg |
| 30.0 Kg | 5KG,10KG,20KG | 0.005kg |
| 60.0 Kg | 5KG,20KG,40KG | 0.01kg |
| 150.0 Kg | 10KG,50KG,100KG | 0.02kg |
| 1500.0 Kg | 100KG,500KG,1000KG | 0.2 kg |
| 320GM | 100MG,1GM,200GM | 0.001gm |
| 310GM | 100MG,1GM,200GM | 0.001gm |
| 210GM | 100MG,1GM,20GM | 0.01gm |

    - Place the standard weights one by one on the weighing platform/ pans and check the readings. Record the readings in Annexure –I

    - A balance shall be considered satisfactory for use if the variance is within ± 0.1% of the selected standard weight.

    - If the error is exceeding the variance limit, put an “Out of Order” label and arrange for rectifying the same by the Service Engineer. Do not use the balance till it is rectified.

    - Use always-Standard weights and balances that are certified by Weights and Measures    Department.

    - Calibration of Balances shall be done  on every day at the start of the operation and record in  Annexure –I

NOTE:   For 320gm Sartorius and 310gm Mettler Toledo Balances, manual calibration shall not be done, because these balances are provided with auto calibration method. To calibrate these Balances press “CAL” button. Linearity of these balances shall be checked as mentioned in point no.5.11

  - Operation:

    - Adjust the Spirit level at Zero position.

    - Switch on the mains.

    - Press Tare then put the material which is to be weighed.

    - Check and Record the weight.

  - Cleaning:

    - Wipe the body of balance using wet duster followed by dry duster.

  - Precaution:

    - Care must be taken to prevent side hammering on weighing plateform.

    - Do not keep weight more then the capacity of the Balance.

    - While operating Balance care should be taken that weight must not drop on the weighing.

  - Enclosure:

Annexure –I Balance Calibration Record.

Change History Log

SOP Training Record

- Abbreviation:

SOP		:	Standard Operating Procedure

No.		:	Number

IPQA		:	In process quality Assurance

Std wt.		:	Standard weight

Obs wt.		:	Observed wt.